## Slide 1: Microtomes
Brian Cheung
Bernard Hwang
David Yu

## Slide 2: Background
The earliest form of microtomy was the freehand sectioning of fresh or fixed material using a sharp razor.
The first microtome used were manual sectioning instruments made from wood, that were mainly used in Botany for cutting plants.
Modern microtomes are precision instruments that cut uniformly thin sections of a variety of materials for microscopic examination.

## Slide 3: Parts of a Microtome

## Slide 5: Capabilities
Standard microtomes are used to cut paraffin imbedded tissue specimens into thin plane sections.
Sectioned material must be reproducible from only a fraction of a micron to a few microns thick. (Leitz 1512: 1 – 25m)
Movements must be precisely controlled to ensure quality of the sectioned material.
  Thickness
  Reproducibility
  Plane parallelism

## Slide 6: Rotary Microtome Operation
Insert paraffin-embedded sample into sample holder and release spring loaded clamp.
Clamp blade into blade holder.

## Slide 7
Adjust blade holder angle and tighten

## Slide 8
Adjust thickness of sample section.
Set counter dial to zero.
Turn crank (one revolution per section).

## Slide 9
microtome pics\MVI_3845.AVI

## Slide 10: Precision Engineering Applications
Constraint of the Sample Block
Constraint of the Knife Blade
Microtome Knife Material Selection
Translational Guide Rails
Digital Volumetric Imaging

## Slide 11: Constraint of the Sample Block
Embedded sample is held against a flat surface and clamped down.
  Semi-kinematic Constraint With A Line Contact

## Slide 12: Constraint of the Knife Blade
Knife is placed in a retaining platform and secured rigidly in place to the microtome by preloading the opposite side of the knife.
If disposable knives are used, the knife is first held in a fixture; and the fixture is secured to the microtome.

## Slide 13: Alternative Constraint Methods
Newer microtomes may use additional methods to secure the blade.
  Additional Tightening Screws For Additional Point Contacts
  Knife Holder With Magnetic Retaining Points

## Slide 14: Microtome Knife Material Selection
Typical microtome blades are manufactured from high carbon steel or heat-treated tool grade steel.
  Free From Impurities
  Anti-corrosive And Rust Resistant
Profile of steel knives
  A: Strongly Plano Concave
  B: Plano Concave
  C: Wedge Shaped (Wedge Angle 27o)
  D: Plane Shaped (Wedge Angle 45o)

## Slide 15: More on Microtome Knives
Depending on the application other materials are available for blade:
  Tungsten Carbide
  Glass
  Diamond
  Sapphire
  Disposable Blades (Stainless Steel)
Sharpening
  Coarse sharpening
  Fine sharpening
  Stropping
  Performed less with the introduction of disposable knives

## Slide 16: Translational Guide Rails
Guide rails with V-grooves used to limit motion to 1-DOF translation.
  Forward Knife Motion
  Vertical Raising And Lowering Of The Rigidly Held Sample

## Slide 17
Previous Sliding Guides Replaced With Rigid And Play-free Cross-roller Linear Bearings
  Reduces Wear On The Grooves
  Relieves Stress In Bearings
  Support More Load With Increasing Contact Surface
Roller bearings have less tendency to move out of position, or creep
Pictures From References 6 and 7

## Slide 18: Digital Volumetric Imaging
Normally
Microtome → Slices → Microscope Slide → Microscope → Viewing or Photography
Digital Volumetric Imaging (DVI)
Microtome → Slices → Digital Photography → Computer → 3-D Model
Precision in Both a Microtome System and a Microscopic Imaging System
Picture From Reference 10

## Slide 19: Examples of DVI Datasets
Pictures From Reference 10

## Slide 20: Focus and Color Images
CCD Chip → Grayscale Image
  Color Filters
Capturing Images: Time-Dependent
Illumination and Motors → Heat → Expansion → Focus Issues

## Slide 21: Solution to Expansion of Block
Add Refrigeration System
Add Auto Focus System
  Objective Lens on Piezo

## Slide 22: More Precise Motion
Higher Magnification Requires Thinner Sections.
Cross Roller Bearings → Air Bearing Slide
Stepper Motor & Linkages → Linear Motor

## Slide 23: References
Steedman, H.F.; Section Cutting in Microscopy; Blackwell Scientific Publications Ltd, Oxford; England, 1960.
Leitz,Ernst; Leitz 1512/1516 rotary microtomes Instructions; Germany; 1984.
http://www.materials-sectioning.com
Woods and Ellis, Microtome: function and design http://home.primus.com.au/royellis/microt/microt.htm
http://biomicro.sdstate.edu/Hildrethm/Bio545/Techniques/Microtom.htm
http://www.schneeberger.com/
http://www.mellesgriot.com/glossary/wordlist/glossarydetails.asp?wID=136
http://www.theproductfinder.com/component/crorol.htm
Collins, Rodney; Personal Interview; San Francisco General Hospital, March 14, 2003
http://www.meyerinst.com/html/resolve3d/default.htm

## Slide 24: Thank You Very Much!
Any Questions?